Isolated 100% Analog ECG Acquisition System ECE 212P — Analog Circuits Lab ¦ Project Flowchart Design of a 100% analog biopotential acquisition system using NE555 charge pump, discrete INA, RLD, and active filter bank

STAGE 1 — Isolated Power Supply (NE555 Charge Pump)

9V Battery (Patient-isolated) Single-supply source — no mains connection for patient safety

↓

NE555 Timer IC — Astable mode (~15 kHz) Generates square wave; drives Dickson charge-pump network

↓

Dickson Charge Pump 1N4148 diodes + 10 µF film capacitors — produces stable negative rail (−V)

↓

Dual-rail Supply: +V / −V Enables op-amp output to swing below ground — prevents clipping of ECG waveform

Pair 1 Demo Tasks - Verify 15 kHz square wave on the oscilloscope - Probe each charge-pump stage to demonstrate voltage doubling - Show stable −V rail under op-amp load (no drooping) - Explain patient-isolation rationale vs mains-derived supply

STAGE 2 — Biopotential Extraction (INA + RLD)

ECG Electrodes — RA, LA, RL Attached to wrist/forearm; raw signal ≈ 1 mV buried in noise

↓  Differential pair into INA

| Input Buffer A (Op-amp 1) Buffers RA differential input |  | Input Buffer B (Op-amp 2) Buffers LA differential input |
| --- | --- | --- |

↓  Both buffers feed difference stage

| Difference Stage (Op-amp 3) 1% metal-film resistors — applies differential gain; subtracts common-mode noise. | RLD Circuit (Op-amp 4) Senses residual CM voltage Injects 180° correction back to RL electrode → Dramatically improves CMRR, suppresses 50 Hz at source |
| --- | --- |

↓  Amplified ECG output

INA Output — Amplified ECG Heartbeat waveform clearly visible; 50 Hz interference suppressed at source

Pair 2 Demo Tasks - Attach RA, LA, RL electrodes to wrist/forearm and show raw 1 mV signal buried in noise - Show INA gain — amplified heartbeat waveform clearly visible on scope - Toggle RLD off → show 50 Hz noise increase; re-enable → show suppression - Explain INA gain equation and 1% tolerance resistor justification

STAGE 3 — Active Filter Bank (Second TL074)

High-Pass Filter — Sallen-Key, 2nd order, 0.5 Hz Removes electrode drift and DC offset below 0.5 Hz

↓

Twin-T Notch Filter — 50 Hz ≥ 40 dB attenuation of mains interference not removed by RLD

↓

Low-Pass Filter — Sallen-Key, 2nd order, 150 Hz Removes high-frequency noise above diagnostic ECG band

↓

Combined bandwidth: 0.5 Hz – 150 Hz Standard diagnostic ECG bandwidth — all three filters work in sequence

↓  Final cleaned signal

Lab Oscilloscope — Final Output Clean, baseline-stable ECG waveform: P wave, QRS complex, T wave clearly identifiable

Pair 3 Demo Tasks - Connect INA output to filter chain and show the final cleaned ECG output on scope - Inject 50 Hz sine wave at filter input — verify ≥ 40 dB notch attenuation at output - Sweep function generator 0.1–1 Hz — demonstrate HPF 0.5 Hz cutoff - Show final clean ECG with identifiable P, QRS, and T waves - Explain Twin-T component matching and frequency equation

Key Components & Budget

| Component | Role in System |
| --- | --- |
| NE555 Timer IC × 6 | Charge pump oscillator at ~15 kHz (astable mode) |
| TL074 Quad Op-Amp × 3 | INA input buffers, difference stage, RLD, full filter bank |
| 1N4148 Signal Diodes × 20 | Dickson charge-pump network |
| 1% Metal-Film Resistors (~60) | INA gain accuracy; filter component matching |
| Ceramic/Film Capacitors (~40) | Charge pump (10 µF), filter RC networks |
| ECG Electrodes + Snap Cables | Patient signal acquisition (RA, LA, RL) |
| 9V Batteries × 3 + Connectors | Patient-isolated power supply |
| Jumper Wires, Header Pins, Misc | Breadboard wiring and prototyping |