ECE 212P — Analog Circuits Lab

PROJECT PROPOSAL

Design of an Isolated, 100% Analog Biopotential (ECG) Acquisition System using a 555 Timer Charge Pump

Team Members

| S.No. | Name | Roll Number |
| --- | --- | --- |
| 1 | Garaga Karthikeya | IMT2024073 |
| 2 | Nuthigattu Yagneswar Gupta | IMT2024063 |
| 3 | Gonuguntla Lohith Prapoorna Chandra | BT2024260 |
| 4 | Eluri Sri Naga Sairam Kapish | IMT2024027 |
| 5 | Kandagatla Venkata Rohith Sai | IMT2024042 |
| 6 | GVN Mokshagna | BT2024234 |

1. Project Idea

We propose to design and build an analog system that acquires a real human electrocardiogram (ECG) signal — a biopotential of approximately 1 mV — and displays a clean QRS waveform on a laboratory oscilloscope. The entire signal chain, from power generation to final filtering, uses only passive components, diodes and op-amps. No microcontrollers, FPGAs, or ADCs are used anywhere in the design.

Stage 1 — Isolated Power Supply (NE555 Charge Pump)

An NE555 timer IC is configured in astable mode at approximately 15 kHz to drive a Dickson charge-pump network built from 1N4148 diodes and 10 µF film capacitors. This generates a stable −V negative rail from a single, patient-isolated battery source. The negative rail is mandatory: it powers the op-amp signal chain in a true dual-supply configuration, ensuring output waveforms can swing below ground without clipping — a fundamental requirement for biopotential acquisition.

Stage 2 — Biopotential Extraction (Instrumentation Amplifier + RLD)

A TL074 Quad Op-Amp implements a fully discrete three-op-amp Instrumentation Amplifier (INA) with a differential gain. Two of the four op-amps form the input buffer stage; the third implements the difference stage. The fourth op-amp is configured as an active Right Leg Drive (RLD) circuit: it senses residual common-mode voltage and injects a 180° phase-shifted correction signal back into the patient's right leg electrode, actively suppressing 50 Hz mains interference at the source and dramatically improving Common-Mode Rejection Ratio (CMRR). All components are 1% tolerance metal-film resistors to preserve gain accuracy.

Stage 3 — Active Filter Bank

A second TL074 implements the complete filter chain. A 0.5 Hz second-order Sallen-Key High-Pass Filter removes electrode drift and DC offset. A Twin-T Notch Filter centred at exactly 50 Hz provides ≥40 dB attenuation of mains interference not suppressed by the RLD circuit. A 150 Hz second-order Low-Pass Filter removes high-frequency noise. The combined filter chain defines the standard diagnostic ECG bandwidth (0.5 Hz – 150 Hz). The final output is connected directly to a lab oscilloscope for display.

2. Component List & Budget

| Component / Item | Qty |
| --- | --- |
| NE555 Timer IC | 6 |
| TL074 Quad Op-Amp | 3 |
| ECG Electrodes + Snap Cables | 1 set |
| 1% Metal-Film Resistors (assorted) | ~60 pcs |
| Ceramic / Film Capacitors (assorted) | ~40 pcs |
| 1N4148 Signal Diodes | 20 |
| Battery + Snap Connectors | 3 sets |
| Jumper Wires, Header Pins, Misc | Lot |

3. Demonstration Plan (April 27–30, 2026).

Pair 1 — Power Section (NE555 Charge Pump)

- Configure the NE555 in astable mode and verify a square wave on the oscilloscope.

- Show the diode-capacitor network charging up: probe each stage to demonstrate voltage doubling.

- Demonstrate the stable a negative voltage rail under load (op-amps connected), confirming the supply can sink the required current without drooping.

- Explain the role of isolation: why a dedicated battery and charge-pump are used instead of a mains-derived supply for patient safety.

Pair 2 — Extraction Section (INA + RLD)

- Attach ECG electrodes (RA, LA, RL) to a team member's wrist/forearm.

- Demonstrate raw differential input on the scope (~1 mV, buried in noise).

- Show the amplified INA output gain — the heartbeat waveform should be clearly visible.

- Demonstrate RLD effectiveness: disable the RLD op-amp and show increased 50 Hz noise; re-enable it and show suppression.

- Explain the INA gain equation and resistor selection; justify 1% tolerance components.

Pair 3 — Filter Section (HPF + Notch + LPF)

- Connect INA output to the filter chain and show the final cleaned output on the scope.

- Demonstrate the 50 Hz notch: inject a 50 Hz sine wave from a function generator at the filter input and show ≥40 dB attenuation at the output.

- Demonstrate HPF roll-off: sweep a function generator from 0.1 Hz to 1 Hz and show the 0.5 Hz cut-off.

- Show the final clean, baseline-stable ECG waveform with clearly identifiable P, QRS, and T waves on the oscilloscope.

- Explain Twin-T notch component matching and frequency equation.